## Slide 1: Chapter 28 High-Performance Liquid Chromatography
Mobile Phase: Liquid
Stationary Phase Separation Mechanism
- Solid Adsorption
- Liquid Layer Partition
- Ion exchange resin Ion exchange
- Microporous beads Size Exclusion
- Chemically modified resin Affinity

## Slide 2: HPLC Advantages vs GC
Not limited by sample volatility or thermal stability
Two interacting phases
Room temperature analysis
Ease of sample recovery

## Slide 3: Instrumentation
Solvent Reservoirs
Pump
Sample Injector
Column(s)
Detector
Data System

## Slide 5: Mobile Phase Reservoirs
Inert container with inert lines leading to the pump are required.
Reservoir filters (2-10 mm) at reservoir end of solvent delivery lines
Degassed solvent
- Vacuum filtration
- Sparge with inert gas (N2 or He)
- Ultrasonic under vacuum
Elevate above pumps

## Slide 6
Isocratic elution: A separation that employs a single solvent or solvent mixture of constant composition.
Gradient elution: Here two or more solvent systems that differ significantly in polarity are employed. After elution is begun; the ratio of the solvents is varied in a programmed way, sometimes continuously and sometimes in a series of steps. Separation efficiency is greatly enhanced by gradient elution.

## Slide 7: HPLC Pump Criteria
Constructed of materials inert toward solvents to be used
Deliver high volumes (flow rates) of solvent (to 10 mL/min)
Deliver precise and accurate flow (<0.5% variation)
Deliver high pressure (to 6000 psi)
Deliver pulse free flow
Have low pump-head volume
Be reliable

## Slide 8: HPLC Pumps: Types
Reciprocating pumps
Syringe pumps
Constant pressure pumps

## Slide 9: Reciprocating Pumps
One, two, or three pump heads
- more heads, less pulse
Small head volumes (50 to 250 mL)
Short piston stroke
Inert pistons (generally sapphire)
Continuous use (no refill time)
Pulse dampeners

## Slide 11: Syringe Pumps
Constant flow rate pump
Non-pulsating flow
Low flow rates (1 to 100 mL/min)
Isocratic flow only
Refill required when reservoir (~50mL) expended

## Slide 12: Constant Pressure Pump
Constant pressure pump, not constant flow
Can deliver high pressures
Stable flow during delivery stroke
Stop flow on refill stroke
Low cost

## Slide 13: Sample Introduction
Valve-type injectors
- Six port fixed volume Rheodyne
reproducible injection volumes
variable loop size
easy to use, reliable
- Six port variable volume Waters
variable injection volumes without loop change increased maintenance, operator skill required more expensive

## Slide 15: Auto Injectors
Continuous injections operator free
Comparable precision and accuracy to manual
Much more expensive initially
Much more convenient Up 100 samples and standards with microprocessor control

## Slide 16
Liquid-Chromatographic Columns
Liquid-chromatographic columns are ordinarily constructed from smooth-bore stainless steel tubing, although heavy-walled glass tubing is occasionally encountered. The latter is restricted to pressures that are lower than about 600 psi.

## Slide 17
Analytical Columns
Liquid-chromatographic columns range in length from 10 to 30 cm. Normally, the columns are straight, with added length, where needed, being gained by coupling two or more columns together. The inside diameter of liquid columns is often 4 to 10 mm; the most common particle size of packings is 5 or 10 m. The most common column currently in use is one that is 25 cm in length, 4.6 mm inside diameter, and packed with 5 m particles. Columns of this type contain 40,000 to 60,000 plates/meter.

## Slide 18
Guard Columns
A guard column is introduced before the analytical column to increase the life of the analytical column by removing not only particulate matter and contaminants from the solvents but also sample components that bind irreversibly to the stationary phase. The guard column serves to saturate the mobile phase with the stationary phase so that losses of this solvent from the analytical column are minimized. The composition of the guard-column packing is similar to that of the analytical column; the particle size is usually larger. When the guard column has become contaminated, it is repacked or discarded and replaced with a new one.

## Slide 19
Detectors: Unlike gas chromatography, liquid chromatography has no detectors that are as universally applicable and as reliable as the flame ionization and thermal conductivity detectors. A major challenge in the development of liquid chromatography has been in detector improvement.
Types of Detectors: Liquid chromatographic detectors are of two basic types. Bulk property detectors respond to a mobile-phase bulk property, such as refractive index, dielectric constant, or density. In contrast, solute property detectors respond to some property of solutes, such as UV absorbance, fluorescence, or diffusion current, that is not possessed by the mobile phase.

## Slide 21
Absorbance Detectors: Is a Z-shaped, flow-through cell for absorbance measurements on eluents from a chromatographic column. Many absorbance detectors are double-beam devices in which one beam passes through the eluent cell and the other through a filter to reduce its intensity.
Ultraviolet Absorbance Detectors with Filters: The simplest UV absorption detectors are filter photometers with a mercury lamp as the source. Most commonly the intense line at 254 nm is isolated by filters. Deuterium or tungsten filament sources with interference filters also provide a simple means of detecting absorbing species.

## Slide 23
UV Absorbance Detector with Monochromator: There are detectors that consist of a scanning spectrophotometer with grating optics. Some are limited to uv radiation; others encompass both uv and visible radiation. The most powerful uv spectrophotometric detectors are diode-array instruments.
Infrared Absorbance Detectors: Two types of infrared detectors are offered commercially. The range of the first instrument is from 2.5 to 14.5 m or 4000 to 690 cm-1. The second type of infrared detector is based upon Fourier transform instruments.

## Slide 24
Fluorescence Detectors: Fluorescence is observed by a photoelectric detector located at 90 deg to the excitation beam. The simplest detectors employ a mercury excitation source and one or more filters to isolate a band of emitted radiation. More sophisticated instruments are based upon a Xenon source and employ a grating monochromator to isolate the fluorescent radiation. An inherent advantage of fluorescence methods is their high sensitivity, which is typically greater by more than an order of magnitude than most absorbance procedures.

## Slide 25
Refractive-Index Detectors: Refractive-index detectors have the significant advantage of responding to nearly all solutes. That is they are general detectors analogous to flame or thermal conductivity detectors in gas chromatography. In addition, they are reliable and unaffected by flow rate. They are, however, highly temperature sensitive and must be maintained at a constant temperature to a few thousandths of a degree centigrade. Furthermore, they are not as sensitive as most other types of detectors.

## Slide 26
Electrochemical Detectors: These devices are based upon amperometry, polarography, coulometry, and conductometry. They appear to offer advantages, in many instances, of high sensitivity, simplicity, convenience, and widespread applicability.

## Slide 28
Mass Spectrometric Detectors: A problem in coupling liquid chromatography with mass spectrometry is the enormous mismatch between the relatively large solvent volumes and the vacuum requirements. Several interfaces have been developed for solving this problem. In a first type the eluent from the column is split, with only a tiny fraction being introduced directly into the mass spectrometer. In a second type of interface the effluent is deposited on a continuous, moving-belt or moving-wire that transports the solvent and analyte to a heated chamber for removal of the former by volatilization.

## Slide 30: PARTITION CHROMATOGRAPHY
Partition chromatography can be subdivided into
(i) liquid-liquid chromatography and
(ii) bonded-phase chromatography.
With liquid-liquid, a liquid stationary phase is retained on the surface of the packing by physical adsorption.
With bonded-phase, the stationary phase is bonded chemically to the support surfaces.

## Slide 31
PARTITION CHROMATOGRAPHY
Early partition chromatography was the liquid-liquid type; now the bonded-phase method has become predominate because of certain disadvantages of liquid-liquid systems.
One of these disadvantages is the loss of stationary phase by dissolution in the mobile phase, which requires periodic recoating of the support particles.
Furthermore, stationary-phase solubility problems prohibit the use of liquid-phase packings for gradient elution.

## Slide 32
Columns for Bonded-Phase Chromatography
The supports for the majority of bonded-phase packings for partition chromatography are prepared from rigid silica, or silica-based, compositions. These solids are formed as uniform, porous, mechanically sturdy particles commonly having diameters of 3, 5, or 10m. The surface of fully hydrolyzed silica is made up of chemically reactive silanol groups. The most useful bonded-phase coatings are siloxanes formed by reaction of the hydrolyzed surface with an organochlorosilane.

## Slide 36
Reversed-Phase and Normal-Phase Packings Two types of partition chromatography are distinguishable based upon the relative polarities of the mobile and stationary phases.
Liquid chromatography based upon highly polar stationary phases such as water or triethyleneglycol supported on silica or alumina particles; a relatively nonpolar solvent such as hexane or I-propylether then served as the mobile phase. This type of chromatography is referred to as normal-phase chromatography.

## Slide 37
In reversed-phase chromatography, the stationary phase is nonpolar, often a hydrocarbon, and the mobile phase is relatively polar (such as water, methanol, or acetonitrile).
In normal-phase chromatography, the least polar component is eluted first because it is the most soluble in the mobile phase; increasing the polarity of the mobile phase has the effect of decreasing the elution time.
In contrast, in the reversed-phase method, the most polar component appears first, and increasing the mobile phase polarity increases the elution time.

## Slide 42: HPLC Derivatization Methods
Why derivatize?
Enhance detector response
Improve analyte resolution
Improve analyte peak shape
Improve analyte sensitivity
Establish analyte identity
Improve analyte stability during analysis
Change analyte physical properties

## Slide 43: Preparative Chromatography
Separation and isolation of relatively large quantities (> 0.1g) of solute
Similar systems to analytical chromatography - higher flow rates (10 to 200 mL/min) - Large sample loops
- preparative columns - larger dimensions and packing size
- detection not critical - not trace analysis - but non-destructive

## Slide 44: Chromatographic Separations
Open column chromatography
- Very slow analysis
- Poor chromatographic efficiency - Inconvenient sample recovery
HPLC
- More rapid analysis
- High efficiency packing materials - continuous flow-through detection

## Slide 45: Gel Permeation LC
Also known as size exclusion or gel filtration
Separation by effective size in solution
- dependent on molecular size and shape - molecules large enough to be excluded from pores all co-elute at tm
- smaller molecules permeate the sp pores to differing degrees based on relative size and are proportionally retarded.

## Slide 46: GPLC Stationary Phases
Stationary phase controls retention or selection
Two types: rigid cross-linked polymer gels and other polymer gels
Classification by pore size, or range of pore sizes
Small molecules - pores 60 to 100Å
Mixed bed columns for wide size ranges

## Slide 48: GPLC Mobile Phases
Mobile phase not generally used to control retention or selection
Selection primarily based on solubility
- low viscosity solvent
- detector compatible
- column compatible
polstryrene - less polar solvents silica gel - water, wide range of solvents, limited pH

## Slide 49: Ion Exchange Chromatography
Reversible exchange of ionic species between the stationary phase and mobile phase
Ionic species chemically bound to insoluble matrix serves as exchange site (adsorption)

## Slide 50: IEC Separations
Insoluble matrix (M+) and counter ion (E-) as stationary phase
Analyte ion (A-) in mobile phase
M+E- + A- M+A- + E-

## Slide 51: IEC Stationary Phases
Silica based materials
- Pellicular particles
Organic materials
- porous beads
- styrene/divinylbenzene crosslinked co- polymers
- methacrylic acid/divinylbenzene crosslinked co-polymers
Inorganic materials

## Slide 53: IEC Stationary Phases
Functional group addition through surface reactions with appropriate reagents to produce the desired cation or anion exchange resin
Cation exchangers (strong & weak)
- Acid functional groups
Anion exchangers (strong & weak)
- Basic functional groups

## Slide 54: Resin Cross-linkage
Pore size a degree of crosslinkage
High degrees of Crosslinkage:
- increase mechanical strength of the resin
- decrease the degree of swelling
- decrease the permeability of the resin

## Slide 55: IEC Mobile Phases
Aqueous solutions of salt or salts with a small % of organic modifier and perhaps a buffer

## Slide 56: Retention & Selectivity Factors
Size and shape of the solvated molecule
Mobile phase pH
Concentration & type of competing ion
Addition of organic solvent to mobile phase
Column temperature

## Slide 57: Cation M-A+ Bond Strengths
Ce3+ > Al3+ >> Ba2+ > Pb2+ > Ca2+ > Ni2+ > Cu2+ > Mg2+ > UO22+ >> Tl+ > Ag+ > K+ > NH4+ > H+

## Slide 58: Anion M+A- Bond Strengths
Citrate > SCN- > CrO42- > SO42- > NO3- > Br- > CN- > Cl- > HCO3- > HCOO- > OH-
Anion selectivities are not as rigidly defined as those of cations

## Slide 59: Prediction of MxAy Bond Strengths
Factors:
- charge on the solute ion
- size of the solvated ion
- degree of resin crosslinkage
- polarizability of the solute molecule
- ion-exchange capacity of the resin
- functional group on the resin
- degree of interaction between the solute and the resin

## Slide 61: Ion Chromatography
Separation of inorganic cations and anions and low molecular weight water-soluble organic acids and bases
Non-suppressed IC methods
Suppressed IC methods

## Slide 63: Ion Exclusion Chromatography
Separation of ionic solutes from weakly ionized of neutral solutes using strong cation and anion exchange resins

## Slide 65: Factors Affecting Retention in IC
Degree of ionization of the solute
- solute pKa
- eluent pH
- organic modifier content
hydrophobic interactions between solute and sp
molecular size of the solute
degree of X-linkage of the sp
system temperature

## Slide 66: Affinity Chromatography
Separation of proteins and biomolecules through the use of a biologically specific immobilized ligand (sp).
Biologically specific ligand is chemically attached to a gel matrix
pH of the system is controlled to control the strength of the interaction between the ligand and the solute molecule

## Slide 67: Complexation Chromatography
Separation based on the rapid reversible formation of a complex between metal ions and ligands
Suitable ligands are immobilized on sp

## Slide 68: Chiral Chromatography
Separation based on the use of chiral stationary phases which interact to varing degrees with optical isomers of the solute
Systems differ little from conventional HPLC systems
Limited applications for a given sp